## Candidates
Candidate ID | Candidate / manufacturer | Function stratum | Population / purpose | Capabilities reported | Metric(s) available | Device-evidence grade | AI-evidence grade | AI interpretation | Conditions / settings | Main limitations / conflicts | Review location | Source IDs | Source URL
BIO | BIOMONITOR IV / BIOTRONIK | Diagnostic monitoring | Arrhythmia monitoring; not pacing or defibrillation | SmartECG plus AI filtering; remote monitoring; PAC/PVC distinction | False-positive reduction; true-episode preservation (external cited source) | Very low | Very low | AI explicitly identified in source B, only for AF detection; filtering in Home Monitoring Service Center, not proof of onboard AI | AF filter = Very Specific; comparator = without SmartECG | Data on file, no sample size or uncertainty; combined SmartECG + AI effect cannot isolate AI increment | pp.4,9–10; refs pp.15–16 | P; B | https://www.prnewswire.com/news-releases/biotronik-announced-first-implant-of-new-implantable-cardiac-monitor-with-artificial-intelligence-targeting-efficiency-in-monitoring-for-cardiac-arrhythmias-301848638.html
MAV | Micra AV2 / Medtronic | Leadless pacing | Bradycardia; ventricular pacing with atrial mechanical sensing | Leadless; accelerometer atrial sensing; automatic AV synchrony; remote monitoring claimed | Projected longevity; mass; claimed heart-rate operating range | Very low | Not demonstrated | Advanced automatic algorithms described; ML model/training/AI-specific benefit not established | Review: 80–100 beats/min synchrony context; pulse settings and pacing burden NR | Review reports 8–13 years and nearly 16 years; projection not observed survival; source M says tracking upper limit 135 beats/min (different metric) | pp.4–6,9–10 | P; M | https://www.dicardiology.com/content/medtronic-receives-fda-approval-its-next-generation-micra-leadless-pacing-systems
MVR | Micra VR2 / Medtronic | Leadless pacing | Bradycardia; ventricular pacing | Leadless pacing; easier programming; remote monitoring claimed | Projected longevity; mass; battery-life improvement claim | Very low | Not demonstrated | AI label in review; no ML model or AI-specific validation in checked sources | Pacing burden, pulse amplitude/width and impedance NR | Review reports 12 years and nearly 17 years; no clinical head-to-head outcome data for these models extracted | pp.4–6,9–10 | P; M | https://www.dicardiology.com/content/medtronic-receives-fda-approval-its-next-generation-micra-leadless-pacing-systems
QIN | Qinming8631 DR / Qinming Medical | Transvenous dual-chamber pacing | Bradyarrhythmia; SSS or AV block; trial ages 18–85 | Dual-chamber rate response; RF programming described in review | 6-month effective pacing versus Talos DR; FAS and PPS | Low | Not demonstrated | Rate response and programmability are not evidence of ML/AI; cited trial does not test AI | Trial: 0.5 ms assessment pulse width; RA threshold ≤1.5 V, RV ≤1.0 V; see evidence rows | QM7211 in review is ventricular lead, not pacemaker; missing endpoints/attrition; source inconsistencies in difference estimates; short follow-up | pp.6–7,9–10; ref p.18 | P; Q | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
BOS | EMBLEM S-ICD + EMPOWER / Boston Scientific | Modular defibrillation + leadless pacing | Ventricular tachyarrhythmia protection; bradycardia pacing and ATP | S-ICD plus communicating leadless pacing module | No numerical clinical performance result in review or checked initiation report | Very low | Not demonstrated | Device coordination is described; no AI-specific validation identified | ATP settings, shock energy, detection thresholds and follow-up NR | Review mixes component and system properties: outside-heart claim applies to S-ICD, not leadless pacemaker; cannot defibrillate applies to pacemaker, not combined system | p.7,9–10; ref p.17 | P; S | https://www.implantable-device.com/2021/12/13/boston-scientific-starts-trial-of-emblem-s-icd-with-empower-leadless-pacemaker/
AVE | AVEIR DR i2i / Abbott | Dual-chamber leadless pacing | Slow/irregular heart rate; dual-chamber pacing | Implant-to-implant communication; synchronous atrial/ventricular pacing | Implant success; successful AV synchrony; activity-specific mean synchrony | Low | Not demonstrated | i2i communication and adaptive pacing described; no ML model or AI-isolating comparator | First 300 patients; 55 centers; three months; seven postures/walking speeds | Single-arm; manufacturer reporting; exact synchrony threshold and endpoint denominators not stated in review/release; no superiority inference | p.8–10; refs pp.15,17 | P; A | https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints

## Evidence
Evidence ID | Candidate ID | Endpoint / capability | Reported value / wording | Relation | Numeric lower / value | Numeric upper | Unit | Evidence layer | Sample / denominator | Follow-up | Comparator | Settings / conditions | GRADE-lite | Grade reason | Review location / source section | Source ID | Source URL | Interpretation / normalization note
E01 | BIO | Monitoring capability | Efficient and accurate arrhythmia monitoring |  |  |  |  | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.4 §3.2; p.9 §5.1 | P | https://doi.org/10.5121/ijcsit.2024.02101 | No sensitivity, specificity, AUROC or false-alert rate supplied in review
E02 | BIO | False-positive detection reduction | 86% | = | 86 |  | % | Cited-source supplement | NR | NR | Without SmartECG | AF filter Very Specific in Home Monitoring Service Center; AI only for AF | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | B: main text and footnotes | B | https://www.prnewswire.com/news-releases/biotronik-announced-first-implant-of-new-implantable-cardiac-monitor-with-artificial-intelligence-targeting-efficiency-in-monitoring-for-cardiac-arrhythmias-301848638.html | Relative reduction; not 86% specificity; combined SmartECG+AI claim; not printed numerically in review
E03 | BIO | True-episode preservation | 98% | = | 98 |  | % | Cited-source supplement | NR | NR | None reported | Same filter and AI scope as E02 | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | B: main text and footnotes | B | https://www.prnewswire.com/news-releases/biotronik-announced-first-implant-of-new-implantable-cardiac-monitor-with-artificial-intelligence-targeting-efficiency-in-monitoring-for-cardiac-arrhythmias-301848638.html | Preservation claim; do not relabel as clinical sensitivity without validation methods
E04 | MAV | Battery longevity projection | nearly 16 years | approximately | 16 |  | years | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.5 §4.4.1; p.10 §5.3 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Confirmed as median projected longevity in M; assumptions NR; not observed durability
E05 | MAV | Battery life, conflicting specification | 8–13 years | range | 8 | 13 | years | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.5 §4.4.4 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Retained as conflicting review claim; do not average with E04
E06 | MVR | Battery longevity projection | nearly 17 years | approximately | 17 |  | years | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.10 §5.3 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Confirmed as median projected longevity in M; not observed durability
E07 | MVR | Battery life, conflicting specification | 12 years | = | 12 |  | years | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.5 §4.4.4 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Retained as conflicting review claim; do not average with E06
E08 | MAV | Battery-life improvement | approximately 40% improvement | approximately | 40 |  | % | Review extraction | NR | NR | Predecessor Micra generation, unspecified in review | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.10 §5.3 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Relative improvement in projected battery life; not patient outcome
E09 | MAV | Device mass | 1.75 grams | = | 1.75 |  | g | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.10 §5.3 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Shared statement for AV2 and VR2; not an independent clinical endpoint
E10 | MVR | Battery-life improvement | approximately 40% improvement | approximately | 40 |  | % | Review extraction | NR | NR | Predecessor Micra generation, unspecified in review | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.10 §5.3 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Relative improvement in projected battery life; not patient outcome
E11 | MVR | Device mass | 1.75 grams | = | 1.75 |  | g | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.10 §5.3 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Shared statement for AV2 and VR2; not an independent clinical endpoint
E12 | MAV | Heart-rate range for claimed improved AV synchrony | 80–100 bpm | range | 80 | 100 | beats/min | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.6 §4.4.4 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Operating context, not measured % synchrony or tracking upper limit; M separately states upper limit 135 beats/min
E13 | MAV | Projected one-device-for-life proportion | 80% of patients only need one device for life | = | 80 |  | % | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.5 §4.4.1 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Projection, not lifetime follow-up; M states >80% for Micra recipients, not model-specific observed result
E14 | QIN | RF-powered programming | Program control without consuming battery power |  |  |  |  | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.6 §4.5.2 | P | https://doi.org/10.5121/ijcsit.2024.02101 | Review claim; energy consumption not quantified
E15 | QIN | Effective pacing, full analysis set | 81.4% vs 79.5% | = | 81.4 |  | % | Cited-source supplement | 113 Qinming; 112 Talos DR (225 total) | 6 months | Talos DR: 79.5% | 14-center randomized single-blind noninferiority trial; margin −9.5 percentage points; missing values LOCF then non-effective | Low | RCT starts High; −1 missing endpoints/attrition; −1 imprecision/reporting inconsistencies and wide noninferiority margin | Q: Abstract; Methods; Results primary endpoints | Q | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/ | Acceptable capture threshold/sensing composite; not % beats paced. Published difference estimates inconsistent with raw proportions; not propagated
E16 | QIN | Effective pacing, per protocol | 95.4% vs 89.5% | = | 95.4 |  | % | Cited-source supplement | 87 Qinming; 86 Talos DR | 6 months | Talos DR: 89.5% | Same trial; complete-protocol subset | Low | Post-randomization exclusions/missing endpoints; reporting inconsistencies; short outcome horizon | Q: Abstract; Results primary endpoints | Q | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/ | Do not combine PPS and FAS; noninferiority is not superiority
E17 | QIN | Capture-threshold assessment pulse width | 0.5 ms | = | 0.5 |  | ms | Cited-source supplement | NR | NR | None reported | Primary-endpoint methods; long-term threshold allows upward fluctuation ≤20% | Not applicable | Protocol setting, not clinical effect estimate | Q: Methods, Efficacy assessments | Q | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/ | Figure caption separately says 0.4 ms; retain methods value, flag inconsistency
E18 | QIN | Acceptable atrial capture threshold | ≤1.5 V | <= | 1.5 |  | V | Cited-source supplement | NR | NR | None reported | 0.5 ms; output tested at twice measured capture threshold | Not applicable | Protocol acceptance criterion | Q: Methods, Primary end points | Q | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
E19 | QIN | Acceptable ventricular capture threshold | ≤1.0 V | <= | 1 |  | V | Cited-source supplement | NR | NR | None reported | 0.5 ms; output tested at twice measured capture threshold | Not applicable | Protocol acceptance criterion | Q: Methods, Primary end points | Q | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
E20 | BOS | Coordinated pacing/defibrillation capability | S-ICD can trigger leadless pacemaker ATP |  |  |  |  | Review extraction | NR | NR | None reported | NR | Very low | Uncontrolled secondary/manufacturer claim; methods or uncertainty absent | p.7 §4.6.1–4.6.4 | P | https://doi.org/10.5121/ijcsit.2024.02101 | No implant success, ATP efficacy, shock success, adverse-event rate or numeric settings extracted; S confirms trial initiation only
E21 | AVE | Implant success | 98.3% | = | 98.3 |  | % | Review extraction | First 300 enrolled; endpoint-specific denominator NR | 3 months | None reported | Prospective international single-arm; 55 centers; activity analysis includes seven postures/walking speeds | Low | Single-arm prospective study starts Low; short horizon, manufacturer summary; no comparative/AI inference | p.8 §4.7.2 | P; A | https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints | Procedural success; denominator-specific event count not reconstructed from rounded percentage
E22 | AVE | Patients with successful AV synchrony | more than 97% | > | 97 |  | % | Review extraction | First 300 enrolled; endpoint-specific denominator NR | 3 months | None reported | Prospective international single-arm; 55 centers; activity analysis includes seven postures/walking speeds | Low | Single-arm prospective study starts Low; short horizon, manufacturer summary; no comparative/AI inference | p.8 §4.7.2 | P; A | https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints | Proportion of patients meeting a synchrony criterion, not average percentage of synchronous beats; exact criterion NR in review/release
E23 | AVE | Mean AV synchrony in each activity condition | more than 95% | > | 95 |  | % | Review extraction | First 300 enrolled; endpoint-specific denominator NR | 3 months | None reported | Prospective international single-arm; 55 centers; activity analysis includes seven postures/walking speeds | Low | Single-arm prospective study starts Low; short horizon, manufacturer summary; no comparative/AI inference | p.8 §4.7.2 | P; A | https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints | Mean synchrony in each of seven postures/walking speeds; not interchangeable with proportion of successful patients
E24 | AVE | Safety endpoint | Freedom from device or procedure complications endpoint achieved |  |  |  |  | Review extraction | First 300 enrolled; endpoint-specific denominator NR | 3 months | None reported | NR | Very low | Pass/fail manufacturer summary without numeric event rate or uncertainty in review/release | p.8 §4.7.2 | P; A | https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints | Not 100% complication-free; exact safety rate not supplied in these extracted sources
Note on D2: Source: P; p.4 §3.2; p.9 §5.1; https://doi.org/10.5121/ijcsit.2024.02101
Note on D3: Source: B; B: main text and footnotes; https://www.prnewswire.com/news-releases/biotronik-announced-first-implant-of-new-implantable-cardiac-monitor-with-artificial-intelligence-targeting-efficiency-in-monitoring-for-cardiac-arrhythmias-301848638.html
Note on D4: Source: B; B: main text and footnotes; https://www.prnewswire.com/news-releases/biotronik-announced-first-implant-of-new-implantable-cardiac-monitor-with-artificial-intelligence-targeting-efficiency-in-monitoring-for-cardiac-arrhythmias-301848638.html
Note on D5: Source: P; p.5 §4.4.1; p.10 §5.3; https://doi.org/10.5121/ijcsit.2024.02101
Note on D6: Source: P; p.5 §4.4.4; https://doi.org/10.5121/ijcsit.2024.02101
Note on D7: Source: P; p.10 §5.3; https://doi.org/10.5121/ijcsit.2024.02101
Note on D8: Source: P; p.5 §4.4.4; https://doi.org/10.5121/ijcsit.2024.02101
Note on D9: Source: P; p.10 §5.3; https://doi.org/10.5121/ijcsit.2024.02101
Note on D10: Source: P; p.10 §5.3; https://doi.org/10.5121/ijcsit.2024.02101
Note on D11: Source: P; p.10 §5.3; https://doi.org/10.5121/ijcsit.2024.02101
Note on D12: Source: P; p.10 §5.3; https://doi.org/10.5121/ijcsit.2024.02101
Note on D13: Source: P; p.6 §4.4.4; https://doi.org/10.5121/ijcsit.2024.02101
Note on D14: Source: P; p.5 §4.4.1; https://doi.org/10.5121/ijcsit.2024.02101
Note on D15: Source: P; p.6 §4.5.2; https://doi.org/10.5121/ijcsit.2024.02101
Note on D16: Source: Q; Q: Abstract; Methods; Results primary endpoints; https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
Note on D17: Source: Q; Q: Abstract; Results primary endpoints; https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
Note on D18: Source: Q; Q: Methods, Efficacy assessments; https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
Note on D19: Source: Q; Q: Methods, Primary end points; https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
Note on D20: Source: Q; Q: Methods, Primary end points; https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/
Note on D21: Source: P; p.7 §4.6.1–4.6.4; https://doi.org/10.5121/ijcsit.2024.02101
Note on D22: Source: P; A; p.8 §4.7.2; https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints
Note on D23: Source: P; A; p.8 §4.7.2; https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints
Note on D24: Source: P; A; p.8 §4.7.2; https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints
Note on D25: Source: P; A; p.8 §4.7.2; https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints

## Methods
Item | Definition / rule
Purpose | Paper-bounded evidence map with targeted checks of cited sources, not a systematic review or treatment ranking. Extracted 2026-10-28; supplements published before the 2024 review.
Scope | Six candidate rows across five manufacturers: split Micra AV2 and VR2; retain EMBLEM+EMPOWER as a combined system. Talos DR appears only as the Qinming trial comparator.
Read first | Candidate matrix is the overview. Evidence rows are atomic claims; filter by endpoint, evidence layer and grade before sorting numeric values.
Adapted recipe | No candidate drugs or dosing regimens. Dose is not applicable. Device settings are extracted only where reported; pulse amplitudes, shock energy and algorithm thresholds are NR elsewhere.
Evidence layers | Review extraction = claim printed in supplied review. Cited-source supplement = additional numeric detail from its references; never silently attributed to the review.
Units | Percentages remain numeric values on a 0–100 scale (98.3 means 98.3%); years, months, g, beats/min, V and ms. No unsupported conversion of 3 months to 90 days. No derived effect sizes or pooled score.
Relations | Keep >, <=, approximately and range separately. Bounds are not point estimates. Blank numeric cell = nonnumeric/not reported, never zero.
GRADE-lite | Endpoint-specific pragmatic certainty, not formal GRADE or regulatory assessment. High=4, Moderate=3, Low=2, Very low=1; Not demonstrated is absence of extracted AI validation, not evidence of no AI.
Starting certainty | Randomized direct comparative clinical evidence starts High; prospective single-arm evidence starts Low; manufacturer/data-on-file/narrative capability claims start Very low.
Downgrades | One level for each serious concern: bias/missing endpoints, inconsistency, indirectness, imprecision, selective reporting/publication. Floor Very low; no upgrades. Protocol specifications are Not applicable.
Qinming grading | High minus one for missing endpoints/attrition and minus one for imprecision/reporting limitations = Low for six-month effective pacing. The RCT does not establish long-term safety or AI benefit.
AVEIR grading | Low for short-term descriptive implant/synchrony endpoints; no head-to-head comparator. Safety pass/fail claim Very low because the extracted review/release does not supply an event rate.
AI grading | No extracted study isolates an AI-versus-no-AI clinical benefit. BIOMONITOR explicitly uses AI but its numeric claim is combined SmartECG+AI, data on file. Other candidates: Not demonstrated in checked materials.
Candidate grades | Overview grade summarizes the most informative extracted device-performance evidence, not every capability. Use evidence-row grade for the actual claim.
No ranking | Monitoring, bradycardia pacing and defibrillation differ in indication and outcome. Projected years, diagnostic filtering and clinical endpoint percentages must not share a performance leaderboard.
Missing and conflict flags | NR=not reported in extracted source. Conflicting longevity/specifications remain separate rows. AI model architecture, training dataset, external validation and inference location mostly unreported.
Safety | Review lists possible complications without rates for Micra and Boston. No cross-device safety ranking; no absence-of-harm inference from unreported data. Not a clinical recommendation.
Excluded contextual material | Electronic nose network, neural-implant review, microneedle diabetes paper, germline ethics and generic IoT sources lack candidate-specific implant comparison results in this review. BACKBEAT citation is mismatched to AVEIR discussion, not an extra evaluated candidate.
Verification limits | Five linked cited sources checked; no exhaustive primary-study search, current approval-status audit, independent data-on-file audit, or systematic risk-of-bias instrument. AVEIR NEJM search result not used for extra numeric endpoints.

## Issues
Issue ID | Subject | Location | Issue | Treatment
I01 | Micra AV2 | pp.5,10 | 8–13 years versus nearly 16 years | Retain both; M identifies nearly 16 as median projected longevity.
I02 | Micra VR2 | pp.5,10 | 12 years versus nearly 17 years | Retain both; M identifies nearly 17 as median projected longevity.
I03 | Qinming model | pp.6,9–10 | 8631 DR versus QM7211 | Q methods identify QM7211 as ventricular lead. Use Qinming8631 DR as candidate.
I04 | Boston system | p.7; p.10 | Entire system said to stay outside heart; cannot defibrillate | Separate S-ICD properties from intracardiac leadless pacemaker; combined system includes defibrillation.
I05 | Boston reference | p.17 | 2023 citation links to 2021 trial initiation | Use checked article date 2021-12-13; no results inferred.
I06 | AVEIR | pp.8,17 | BACKBEAT reference attached to AVEIR description | Use Abbott release for AVEIR study; do not import BACKBEAT evidence.
I07 | AI label | Device sections | Automatic algorithms/communication called AI without model validation | Keep device evidence distinct from AI-specific evidence.
I08 | Qinming primary study | Q Results and Methods | Reported differences do not consistently match raw proportions; 0.5 ms methods vs 0.4 ms figure caption | Retain group endpoint percentages; do not calculate superiority or propagate inconsistent effect estimates.
I09 | AVEIR endpoints | p.8 | >97% patients successful vs >95% mean synchrony | Separate denominators and estimands; neither is a head-to-head advantage.
I10 | BIOMONITOR numeric claims | B footnote | 86%/98% depend on Very Specific filter; AI AF-only | Not generic specificity/sensitivity; comparator lacks SmartECG, so AI-only increment unresolved.

## Sources
Source ID | Source | URL | Access / verification | Notes
P | Leocadio (2024), supplied review | https://doi.org/10.5121/ijcsit.2024.02101 | Full supplied PDF, 18 pages | Narrative review; p.15 states no data other than public citations
B | BIOTRONIK (2023-06-12), manufacturer release | https://www.prnewswire.com/news-releases/biotronik-announced-first-implant-of-new-implantable-cardiac-monitor-with-artificial-intelligence-targeting-efficiency-in-monitoring-for-cardiac-arrhythmias-301848638.html | Full web text checked | AI applies only to AF detection; data on file; Very Specific filter
M | Wainscot Media (2023-05-01), manufacturer announcement coverage | https://www.dicardiology.com/content/medtronic-receives-fda-approval-its-next-generation-micra-leadless-pacing-systems | Full web text checked | Median projected longevity; not observed 16–17-year follow-up
Q | Xiang et al. (2016), randomized noninferiority trial | https://pmc.ncbi.nlm.nih.gov/articles/PMC5126155/ | Full HTML including methods/results checked | Correct DOI: 10.4103/0366-6999.193443; QM7211 is a ventricular lead
A | Abbott (2023-05-20), manufacturer trial-results release | https://abbott.mediaroom.com/2023-05-20-Abbott-Unveils-Late-Breaking-Data-Showing-the-Worlds-First-Dual-Chamber-Leadless-Pacemaker-Study-Met-All-Three-Primary-Endpoints | Full web text checked | Prospective single-arm study; 300 patients; 55 centers; three months
S | Prutchi (2021-12-13), trial-initiation report | https://www.implantable-device.com/2021/12/13/boston-scientific-starts-trial-of-emblem-s-icd-with-empower-leadless-pacemaker/ | Full web text checked | Review dates it 2023; linked article is dated 2021; no trial results